AtriCure Cryo Module (ACM)

USER’s manual

Model ACM1 – 115 VAC

Model ACM2 – 230 VAC

AtriCure Incorporated

7555 Innovation Way

Mason, Ohio 45040 USA

Customer Service:

1-866-349-2342 (toll 			free)

1-513-755-4100 (phone)

P000669 Rev. J

European Representative:

Herbert Köntges

Köntges SPRL

Avenue Hellevelt 35

B-1180 Brussels Belgium

Tel: +32 2 375 51 63

FAX: +32 2 375 89 06

email: herbert.kontges@skynet.be

Table of Contents

1.	Getting Started	4

1.1.	System Description	5

1.2.	Unpacking	6

1.3.	Warnings and Precautions	6

1.4.	EMC Guidance and Manufacturer’s Declaration	7

1.5.	Responsibility of the Manufacturer	11

2.	The AtriCure Cryo Module (ACM)	12

2.1.	Device Description	12

2.2.	ACM Front Panel – Illustration and Nomenclature	13

2.3.	ACM Rear Panel – Illustration and Nomenclature	19

3.	Installing the ACM	22

3.1.	Transporting of ACM	22

3.2.	Preparing the ACM For Use	22

3.3.	Installing Power Cord	22

3.4.	Installing the Footswitch	22

3.5.	Installing Cylinder Heater Band	23

3.6.	Installing N2O Gas Line Hose	24

3.7.	Installing N2O Exhaust Hose	26

3.8.	Connecting and Disconnecting the Handpiece	27

4.	Instructions For Use	29

4.1.	Powering Up the ACM	29

4.2.	Operating Modes	30

4.3.	Delivering Cryo Energy	32

4.4.	Parameter Entry Mode / Adjusting the System Default Parameter Values	36

5.	Troubleshooting	37

5.1.	No ACM Power / Display Function	37

5.2.	ACM Fault Codes	37

5.3.	ACM Error Codes	38

5.4.	Handpiece Error Codes	40

6.	Symbols Used	41

7.	Technical Specifications	44

7.1.	Mechanical Specifications	44

7.2.	Environmental Specifications	44

7.3.	Electrical Specifications	44

7.4.	Mains Fuses	44

7.5.	Footswitch Specifications	44

7.6.	Equipment Type / Classification	44

8.	Preventive Maintenance and Cleaning of ACM	45

8.1.	Preventive Maintenance	45

8.2.	Cleaning and Disinfecting	46

9.	Servicing of ACM Unit	47

9.1.	Replacement of AC Mains Fuses	47

9.2.	Inspection and Replacement of In-Line N2O Filter	48

9.3.	Replacement of Gas Line Desiccant Filter and Tip Washer	50

9.4.	Other Replacement Components	52

10.	Accessories	53

10.1.	CMH15, Cryo Module Cylinder Heater Band	53

10.2.	CMH22, Cryo Module Cylinder Heater Band	53

10.3.	CMF1, Cryo Module Footswitch	53

11.	Disposal	53

# Getting Started

This manual and the equipment it describes are for use only by qualified medical professionals trained in the particular technique and surgical procedure to be performed.

Caution:  	Federal (USA) law restricts this device to sale by or on the order of a physician.

Please read all information carefully. Failure to properly follow the instructions may lead to serious surgical consequences.

Important:  This manual is designed to provide instructions for use of the AtriCure Cryo Module (ACM) with the AtriCure Cryo Handpiece (CRYO1 and cryoICE) and AtriCure Accessory Devices (CMF1, CMH15, CMH22).  This manual is not a reference to surgical technique.

The AtriCure Cryo Module (ACM) unit is an electro-mechanical and pneumatic cryogenic surgical system that delivers a cryogenic energy source, namely Nitrous Oxide, to create lines of ablation through cardiac tissue.  The ACM is part of a system which includes the nitrous oxide gas, N2O gas line hose, N2O exhaust hose, cylinder heater band, a footswitch, and single-use cryo handpiece with integrated electronics cable / gas hoses.  The system provides controlled lesion forming temperature that is below -40°C, with typical operating ranges between -50°C to -70°C.  The system is also equipped with the ability to tune or set the terminal handpiece temperature.

The ACM is designed to operate only with AtriCure designed and developed handpieces.  Along with the Ablation Activation Switch on the front panel of the ACM, a footswitch can also be used to activate the ablation cycle.  Refer to the handpiece Instruction for Use for a complete description of the indications and uses of these devices.

For the user’s convenience, the AtriCure Cryo Module will be referred to in this User’s Manual as the “ACM”.  The AtriCure Cryo Handpiece will be referred in this User’s Manual as the “Handpiece”.

This User’s Manual provides a description of the ACM, its controls, displays, indicators, and a sequence for its operation with the Handpiece.  This User’s Manual also supplies other information of importance to the user. This manual is intended as a User’s Manual only.  Do not operate the ACM before thoroughly reading this manual.

## System Description

As shown in Figure 1, the system is comprised of the following:

- A:	AtriCure Cryo Module Cylinder Heater Band (CMH15 or CMH22)

- B:	AtriCure Cryo Footswitch (CMF-1)

- C:	ACM N2O Exhaust Hose

- D:	ACM N2O Gas Line Hose Module

- E:	AtriCure Cryo Module (ACM)

- ACM Power Cord (not shown)

- AtriCure Cryo Handpiece with integral cable (not shown)

Accessory devices are described in Section 10.

A

E

D

C

B

Figure 1 – ACM, Footswitch, N2O Gas Line Hose Module, N2O Exhaust Hose, and Cylinder Heater Band

## Unpacking

### ACM Unit

Lift the ACM unit from the box and remove the protective wrapping and side caps.

Note:  The ACM unit is packed upside down for transit purpose.

Note:  Take precautions when lifting the unit as it weighs approx 45 lbs (20 kg).

### Accessories

ACM accessories are shipped separately from the ACM unit.  The accessory box contains the following items:

- ACM Power Cord

- N2O Gas Line Hose

- N2O Exhaust Hose

- ACM Cylinder Heater Band

- ACM Footswitch

It is recommended that the original shipping box and protective wrapping be saved for future storing and/or transporting of the device.

## Warnings and Precautions

The safe and effective use of the cryo device and equipment is highly dependent upon factors under the control of the operator.  There is no substitute for a properly trained operating room staff.  It is important that the operating instructions supplied with the ACM unit be read, understood, and followed before use.

### WARNINGS

Do not operate the ACM unit before thoroughly reading this manual.

Do not use cryo surgical equipment unless properly trained in the specific procedure being undertaken.  This manual and the equipment it describes are for use only by qualified medical professionals trained in the particular technique and surgical procedure to be performed.

Fire Hazard:  Do not use extension cords.

Trip Hazard:  Standard care should be used to reduce the risk of tripping on the Footswitch cable, as well as the N2O exhaust hose.

No modification of this equipment is allowed.

The voltage selector is factory set and should not be changed by the user.  The voltage setting and the fuse rating must be appropriate as identified to prevent ACM malfunction and potential instrument damage.

- Electric Shock Hazard:  Connect the ACM Power Cord to a properly grounded receptacle.  Do not use power plug adapters.

- Electric Shock Hazard:  Do not connect wet accessories to the generator.

Electric Shock Hazard:  Ensure that the Handpiece is correctly connected to the ACM and that no thermocouple wires are exposed from the cable, connector, or the Handpiece.

### PRECAUTIONS

Use only with the AtriCure Handpieces intended for use with the ACM.

Do not activate the ACM unit until the Handpiece is properly positioned at the ablation site.

The system status indicators and displays are important safety features.  Do not obstruct either the ablation or the system status indicators.

Do not remove the ACM cover, except for the in-line gas filter access port, as there is a potential for electrical shock.  Refer to authorized personnel for service.

Use only the Footswitch provided with the ACM.

The Power Cord of the ACM must be connected to a properly grounded receptacle.  Extension cords and/or adapter plugs must not be used.

Do not contact AtriCure Handpieces with a RF device.

Compressed Air Hazard: Do not operate N2O tanks with a pressure greater than 1000 PSIG (6900 kPa).

## EMC Guidance and Manufacturer’s Declaration

### Electromagnetic Requirements

#### The AtriCure Cryo Module (ACM) has been tested and found to comply with the limits for medical devices in IEC 60601-1-2:2007.  These limits are designed to provide reasonable protection against harmful interference in a typical medical installation.

#### The ACM can radiate radio frequency energy and, if not installed and used in accordance with the instructions, may cause harmful interference to other devices in the vicinity.

#### Portable and mobile RF communications equipment can also affect ACM performance and care should be taken to minimize such interference. However, there is no guarantee that interference will not occur in a particular installation.

#### If the ACM does cause harmful interference to other devices, which can be determined by turning the ACM off and on, the user is encouraged to try to correct the interference by one or more of the following measures:

- Reorient or relocate the receiving device.

- Increase the separation between the ACM and the other devices.

- Connect the ACM into an outlet on a circuit different from that to which the other device(s) are connected.

- Contact the AtriCure service representative for help.

### Electromagnetic Emissions

| Guidance and manufacturer’s declaration – electromagnetic emissions |  |  |
| --- | --- | --- |
| The AtriCure Cryo Module (ACM) is intended for use in the electromagnetic environment specified below. The customer or the user of the ACM unit should assure that it is used in such an environment. |  |  |
| Emissions test | Compliance | Electromagnetic environment – guidance |
| RF emissions CISPR 11 | Group 1 | The ACM unit uses RF energy only for its internal function. Therefore, its RF emissions are very low and are not likely to cause any interference in nearby electronic equipment. |
| RF emissions |  |  |
| CISPR 11 | Class A | The ACM unit is suitable for use in all establishments other than domestic and those directly connected to the public low-voltage power supply network that supplies buildings used for domestic purposes. |
| Harmonic emissions IEC 61000-3-2 | Class A |  |
| Voltage fluctuations/ flicker emissions IEC 61000-3-3 | Complies |  |

### Electromagnetic Immunity

| Guidance and manufacturer’s declaration – electromagnetic immunity |  |  |  |  |
| --- | --- | --- | --- | --- |
| The AtriCure Cryo Module (ACM) is intended for use in the electromagnetic environment specified below. The customer or the user of the ACM unit should assure that it is used in such an environment. |  |  |  |  |
| IMMUNITY test |  | IEC 60601 test level | Compliance level | Electromagnetic environment – guidance |
| Electrostatic discharge (ESD) IEC 61000-4-2 |  | ± 6 kV contact ± 8 kV air | ± 6 kV contact ± 8 kV air | Floors should be wood, concrete, or ceramic tile. If floors are covered with synthetic material, the relative humidity should be at least 30 %. |
| Electrical fast transient/burst IEC 61000-4-4 |  | ± 2 kV for power supply lines ± 1 kV for input/output lines | ± 2 kV for power supply Lines ± 1 kV for input/output lines | Mains power quality should be that of a typical commercial or hospital environment. |
| Surge IEC 61000-4-5 |  | ± 1 kV line(s) to line(s) ± 2 kV line(s) to earth | ± 1 kV differential mode ± 2 kV common mode | Mains power quality should be that of a typical commercial or hospital environment. |
| Voltage dips, short interruptions and voltage variations on power supply input lines IEC 61000-4-11 |  | <5 % UT (>95 % dip in UT) for 0,5 cycle 40 % UT (60 % dip in UT) for 5 cycles 70 % UT (30 % dip in UT) for 25 cycles <5 % UT (>95 % dip in UT) for 5 s | <5 % UT (>95 % dip in UT) for 0,5 cycle 40 % UT (60 % dip in UT) for 5 cycles 70 % UT (30 % dip in UT) for 25 cycles <5 % UT (>95 % dip in UT) for 5 s | Mains power quality should be that of a typical commercial or hospital environment. If the user of the ACM unit requires continued operation during power mains interruptions, it is recommended that the ACM unit be powered from an uninterruptible power supply or a battery. |
| Power frequency (50/60 Hz) magnetic field IEC 61000-4-8 |  | 3 A/m | 3 A/m | Power frequency magnetic fields should be at levels characteristic of a typical location in a typical commercial or hospital environment. |
| NOTE: | UT is the a.c. mains voltage prior to application of the test level. |  |  |  |

### EMC Guidance and Manufacturer’s Declaration

| Guidance and manufacturer’s declaration – electromagnetic immunity |  |  |  |
| --- | --- | --- | --- |
| The AtriCure Cryo Module (ACM) is intended for use in the electromagnetic environment specified below. The customer or the user of the ACM should assure that it is used in such an environment. |  |  |  |
| IMMUNITY test | IEC 60601 TEST LEVEL | Compliance level | Electromagnetic environment – guidance |
| Conducted RF IEC 61000-4-6 | 3 Vrms 150 kHz to 80 MHz | 3 Vrms | Portable and mobile RF communications equipment should be used no closer to any part of the ACM, including cables, than the recommended separation distance calculated from the equation applicable to the frequency of the transmitter. Recommended separation distance d = 1.2 √P |
| Radiated RF IEC 61000-4-3 | 3 V/m 80 MHz to 2.5 GHz | 3 V/m | d = 1.2 √P 80 MHz to 800 MHz d = 2.3 √P 800 MHz to 2.5 GHz where P is the maximum output power rating of the transmitter in watts (W) according to the transmitter manufacturer and d is the recommended separation distance in meters (m). Field strengths from fixed RF transmitters, as determined by an electromagnetic site survey,a should be less than the compliance level in each frequency range.b Interference may occur in the vicinity of equipment marked with the following symbol: |
| NOTE 1: At 80 MHz and 800 MHz, the higher frequency range applies. NOTE 2: These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects, and people. |  |  |  |
| a Field strengths from fixed transmitters, such as base stations for radio (cellular/cordless) telephones and land mobile radios, amateur radio, AM and FM radio broadcast and TV broadcast cannot be predicted theoretically with accuracy. To assess the electromagnetic environment due to fixed RF transmitters, an electromagnetic site survey should be considered. If the measured field strength in the location in which the ACM is used exceeds the applicable RF compliance level above, the ACM should be observed to verify normal operation. If abnormal performance is observed, additional measures may be necessary, such as re-orienting or relocating the ACM. b Over the frequency range 150 kHz to 80 MHz, field strengths should be less than 3 V/m. |  |  |  |

### Recommended Separation Distance

| Recommended separation distances between portable and mobile RF communications equipment and the AtriCure Cryo Module |  |  |  |
| --- | --- | --- | --- |
| The AtriCure Cryo Module (ACM) is intended for use in an electromagnetic environment in which radiated RF disturbances are controlled. The customer or the user of the ACM can help prevent electromagnetic interference by maintaining a minimum distance between portable and mobile RF communications equipment (transmitters) and the ACM as recommended below, according to the maximum output power of the communications equipment. |  |  |  |
| Rated maximum output power of transmitter W | Separation distance according to frequency of transmitter m |  |  |
|  | 150 kHz to 80 MHz d = 1.2 √P | 80 MHz to 800 MHz d = 1.2 √P | 800 MHz to 2.5 GHz d = 2.3 √P |
| 0.01 | 0.12 | 0.12 | 0.23 |
| 0.1 | 0.38 | 0.38 | 0.73 |
| 1 | 1.2 | 1.2 | 2.3 |
| 10 | 3.8 | 3.8 | 7.3 |
| 100 | 12 | 12 | 23 |
| For transmitters rated at a maximum output power not listed above, the recommended separation distance d in meters (m) can be estimated using the equation applicable to the frequency of the transmitter, where P is the maximum output power rating of the transmitter in watts (W) according to the transmitter manufacturer. NOTE 1: At 80 MHz and 800 MHz, the separation distance for the higher frequency range applies. NOTE 2: These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects, and people. |  |  |  |

## Responsibility of the Manufacturer

AtriCure is responsible for safety, reliability, and performance of the equipment only if:

Installation procedures in this manual are followed.

Persons authorized by AtriCure carry out modifications or repairs.

The electrical installation of the relevant room complies with local codes and regulatory requirements such as IEC and BSI.

The equipment is used in accordance with the AtriCure User’s Manual.

# The AtriCure Cryo Module (ACM)

This section provides a detailed description of the ACM including its function and operating features.

## Device Description

- The AtriCure Cryo Module (ACM) unit is an electro-mechanical and pneumatic cryogenic surgical system that delivers a cryogenic source, Nitrous Oxide, to create lines of ablation through cardiac tissue.

- The ACM is part of a system which includes the nitrous oxide gas and its cylinder, N2O gas line hose, N2O exhaust hose, cylinder heater band, a footswitch, system power cord, and single-use cryo handpiece with integrated electronics cable / gas hoses.

- The system provides a controlled lesion forming temperature that is below -40°C, with typical operating ranges between -50°C to -70°C for the treatment of cardiac arrhythmias, creating an inflammatory response (cryonecrosis) that blocks the cardiac electrical conduction pathway.

- The system is also equipped with the ability to tune or set the terminal handpiece temperature.

- Along with the Ablation Activation Switch on the front panel of the ACM, the ACM Footswitch can also be used as the input device to activate and terminate the cryo ablation cycle.

- The presence of the ACM Footswitch is not a requirement for the ACM unit to operate and is made available as a secondary means of interfacing with the ACM unit, and to provide control of the ablation cycle to the surgeon.

- The ACM is designed to operate only with AtriCure designed and developed handpieces.

- Refer to the Handpiece Instruction for Use for complete description of the indications and uses of these devices.

## ACM Front Panel – Illustration and Nomenclature

An illustration of the ACM front panel is shown in Figure 2, below.

1

5

7

12

11

6

3

8

11

2

4

10

9

Figure 2 – ACM Front Panel

| 1. Ablation Activation Switch | 7. Ablation Timer Increase Button |
| --- | --- |
| 2. Ablation Status Indicator | 8. Handpiece Probe Temperature |
| 3. N2O Gas Gauge Display | 9. Handpiece Thermocouple Port |
| 4. System Status Indicator | 10. Handpiece Electronic Interface Port |
| 5. Ablation Timer Decrease Button | 11. Handpiece Gas Connection Port |
| 6. Ablation Timer Display | 12. Handpiece Probe Temperature Control |

### Ablation Status Indicators

There are three indicators on the front panel of the ACM:  READY, FREEZE, and DEFROST.  These indicators are described below.

1

2

3

Figure 3 – ACM Ablation Status Indicators

| Indicator | Symbol | Description |
| --- | --- | --- |
| 1 |  | READY – Green LED This status indicates the ACM is in stand-by mode and is ready to start the cryo ablation cycle. |
| 2 |  | FREEZE – Blue LED This status indicates the ACM is in cryo freeze state and the N2O gas is allowed to flow through the Handpiece causing drop in probe temperature. |
| 3 |  | DEFROST – Amber LED Constant: This status indicates the ACM is in defrost state, N2O gas flow is ceased, Handpiece is back pressurized, and the Handpiece probe is warming back up to the ambient temperature. Blinking: Indicates the ACM is now ready to transition to the next state. Upon activation of either the Ablation Activation Switch or the Footswitch, the ACM unit will transition into the next state depending on the ACM model as described below. International – ACM - Transitions to Ready state only US – ACM - Press and release – Freeze state - Press and hold – Ready state |

### System Status Indicators

There are six system status indicators on the front panel of the ACM:  SYSTEM POWER, SYSTEM FAULT, MAINTENANCE NEEDED, LOW CYLINDER PRESSURE, CYLINDER HEATER BAND ACTIVE, and FOOTSWITCH ACTIVE.  These indicators are described below.

1

2

3

4

6

5

Figure 4 – ACM Status Indicators

| Indicator | Symbol | Description |
| --- | --- | --- |
| 1 |  | SYSTEM POWER – Green LED This status indicates system logic power for the ACM is available. |
| 2 |  | SYSTEM FAULT – Red LED This status indicates ACM has encountered a system fault condition and has halted all operation. The ACM unit will not operate until the fault condition has been addressed. |
| 3 |  | MAINTENANCE NEEDED – Amber LED This status indicates the ACM system is still in normal, functional, and safe operating mode. However, maintenance is being requested to address such items as replacing in-line N2O gas filter to optimize system performance. |
| 4 |  | lOW CYLINDER PRESSURE – Amber LED This status indicates the N2O cylinder gas pressure into the ACM unit is below (initial system power up) or is dropping below (during usage) the optimal operating level. Two LED indication modes are available. Initial Power Up: Constant: Indicates the cylinder pressure is still below the optimal operating level. Blinking: Indicates cylinder valve is closed, or if the valve is open, the cylinder pressure is well below the operating level or is empty of N2O gas. - In both of these cases, the cylinder heater band will be active to raise the cylinder pressure to its recommended operating level as long as the cylinder valve is in the open position. - The cylinder heater band will not be active if the cylinder valve is in the CLOSED position. - The indicator LED will extinguish when the cylinder pressure has reached its optimal operating level. During Usage: Constant: Indicates the cylinder pressure has dropped below the optimal operating level, and the cylinder heater band will be activated. The indicator LED will extinguish when the cylinder pressure has risen back to the optimal operating level. Blinking: Indicates the cylinder pressure is well below the recommended operating level, and the cylinder has less than 10 minutes of functional gas left. |
| 5 |  | CYLINDER HEATER BAND ACTIVE – Amber LED This status indicates ACM unit has energized the N2O cylinder heater band and it is active. The indicator LED will extinguish when the cylinder heater band is turned off by the ACM unit (when optimal cylinder pressure is reached). |
| 6 |  | FOOTSWITCH ACTIVE – Blue LED This status indicates the ACM footswitch pedal is depressed and is in activate state. The indicator LED will extinguish when the footswitch pedal is released. |

### Front Panel Functions

| Item | Description |
| --- | --- |
|  | HANDPIECE THERMOCOUPLE Receptacle This 2-pin receptacle accepts the AtriCure Handpiece thermocouple connectors. This connection is patient-isolated. |
|  | HANDPIECE ELECTRONIC INTERFACE Receptacle This 14-pin receptacle is reserved for future release of AtriCure Handpiece. This connection is patient-isolated. |
|  | HANDPIECE PNEUMATIC Receptacle This 2 port pneumatic receptacle provides the N2O interface with the AtriCure Handpiece. The ports are sized differently to prevent cross-coupling of the connections. |
|  | HANDPIECE PROBE TEMPERATURE ADJUST Knob This control knob provides the operator with the ability to adjust the terminal temperature of the handpiece probe. Turning the knob in the counter-clockwise rotation lowers the probe temperature, while turning the knob in clockwise rotation raises the probe temperature. In normal operation the knob will be adjusted to provide maximum negative temperature. |

### Front Panel Displays

#### N2O Gas Gauge

- This display indicates, in minutes, the approximate amount of useable N2O gas remaining in the cylinder.  As the gas is being consumed, this value will adjust based on the actual gas flow time and the cylinder pressure.  As a result, the displayed value may decrease quicker than anticipated.

- This time value should not be taken as an absolute measurement, but only as a reference in gauging the amount of gas remaining in the cylinder.

- The display value will start to blink when there is approximately 10 minutes or less of anticipated useable gas left in the cylinder.

NOTE:	Replacement of the N2O cylinder at this point is highly recommended.

NOTE:	Along with the blinking gas gauge, the Low Cylinder Pressure indicator LED may also be blinking indicative of empty or near empty N2O cylinder.

- Typical usage time for a new cylinder is between 30 to 40 minutes depending on the type and size of the cylinder when used with the ACM cylinder heater band.  The usage time will be less when the cylinder heater band is not used.

- The gas gauge value can be reset upon replacement of the N2O cylinder by pressing down on the Timer Increase and Timer Decrease key simultaneously and releasing.

- Secondary function of the gas gauge is to display the N2O cylinder gas pressure during the ablation cycle.  This provides the user with additional real-time information regarding the gas status within the cylinder during system operation in psi x10.

#### Ablation Timer

- This display indicates the desired ablation duration in units of seconds.

- The default duration time is 120 seconds per ablation run.

- This timer value can be adjusted prior to, or during the ablation cycle, by using either the Timer Increase or Timer Decrease Button.

- Each key press of the Timer Increase Button or Timer Decrease Button will modify the target ablation time value by 10 seconds, either adding or subtracting from its present value.

- At the termination of the freeze cycle (Defrost mode), the display will show total time in freeze state.  A full 120 second run will result in the display showing 120 as an example.  An aborted freeze cycle will show the number of seconds the system operated in the Freeze mode.

#### Handpiece Probe Temperature

This display indicates the current handpiece probe temperature in Celsius, and is continuously updated each second.

## ACM Rear Panel – Illustration and Nomenclature

An illustration of the ACM rear panel is shown in Figure 5, below.

1

5

31

2

4 3

6

8

7

Figure 5 – ACM Rear Panel

| 1. N2O Exhaust Port | 5. Equipotential Terminal |
| --- | --- |
| 2. System Pressure Relief Knob | 6. Cylinder Heater Band Receptacle |
| 3. N2O Inlet Port | 7. AC Power Entry Module |
| 4. Footswitch Receptacle | 8. Input Voltage Selector Switch |

Note:  Input Voltage Level Switch is factory set and is covered by unit model label.  The switch and the label should not be tampered with or changed except by authorized AtriCure personnel.

### Rear Panel Functions

| Item | Description |
| --- | --- |
|  | Power Entry Module – This module contains both the On/OFF switch and the wall receptacle mains fuses. Fuse Box – The Fuse Box contains fuses selected for the input voltage. See Technical Specifications in Section 7 of this manual. |
|  | Cylinder Heater Band Receptacle – This receptacle accepts the ACM Cylinder Heater Band connector. It provides power for the heating element embedded in the heater band, as well as transmitting back to the ACM unit the N2O cylinder wall temperature. |
|  | Footswitch Receptacle – This receptacle accepts the ACM Footswitch connector. The single momentary actuation pedal provides for the activation and termination of the cryo ablation cycle. |
|  | N2O Inlet Receptacle – This receptacle accepts the N2O gas line hose with the quick connect / disconnect interface used to transfer gas from the N2O gas cylinder to the ACM unit to be utilized by the Handpiece. |
|  | N2O Exhaust Receptacle – This ½ inch (1.25cm) receptacle accepts the 50 ft (15 m) flexible tubing to direct the expelled N2O gas in a manner appropriate with local and hospital regulation. |
|  | System Pressure Relief Knob – This knob, when pulled, will depressurize and expel trapped N2O gas within the ACM unit. Be sure the N2O cylinder valve is CLOSED prior to performing this task. |
|  | Equipotential Terminal – This terminal provides a means of securely linking the earth ground of the AtriCure ACM unit to other grounded equipment. This terminal is accessible during any position of normal use and minimizes risk of accidental disconnection. The conductor is detached without use of a tool. This terminal is not to be used for a protective earth connection. |

# Installing the ACM

Inspect the ACM system for any signs of physical damage to the front and rear panel, N2O gas couplers, connector receptacles, chassis or cover.

NOTE:	If any physical damage is found, DO NOT use THE UNIT.   Contact AtriCure for a replacement.

All returns must be approved by AtriCure.

## Transporting of ACM

The use of handles is highly recommended in lifting, carrying, or moving of the ACM unit.  The ACM weighs approximately 45 lbs (20 kg) so care should be taken when transporting the ACM unit.

## Preparing the ACM For Use

### The ACM unit may be placed on a cart or on any sturdy table or platform.  It is recommended that carts have conductive wheels.  Refer to hospital procedures or local codes for detailed information.

### The ACM unit may be placed together or stacked with other AtriCure capital equipment such as the ASU, ASB, and the ORLab.  However, care must be taken when stacking or placing the ACM unit adjacent to other medical equipment and unit performance should be observed to verify normal operation in the configuration in which it will be used.

### Provide at least four to six inches of space around the sides of the ACM for convection cooling.  Under continuous use for extended periods of time, it is normal for the top panel to be warm to the touch.

## Installing Power Cord

The ACM unit is shipped with an approved hospital grade power cord.

Insert the power cord into the power entry module of the ACM unit as shown in Figure 7, and then into a grounded hospital AC power receptacle.

NOTE:	ACM power cord is shipped with the accessories and found in the accessory package.

NOTE:	Do not use extension cords or three-prong to two-prong adapters.  The Power Cord assembly should be periodically checked for damaged insulation or connectors.

## Installing the Footswitch

### Inspect the Footswitch

Inspect the Footswitch for any signs of physical damage to the cable and connector.  If physical damage is found or the Footswitch does not perform within specification, contact AtriCure for replacement.

### Connecting and Disconnecting the Footswitch

With the connector alignment arrow in the 12 o’clock position, push the Footswitch Connector into the Footswitch Receptacle on the rear panel of the ACM as shown in Figure 7.  To disconnect, firmly grasp the connector body and pull out from the Footswitch Receptacle.

NOTE:	Typically, you will connect the Footswitch to the ACM prior to powering up the unit.  However, the Footswitch may be connected to the ACM when the unit is powered up and is in READY operating mode (see Section 4 regarding the READY mode).

### Preparing the Footswitch for Use

The Footswitch should be placed on a flat floor.  It is recommended that the area near the Footswitch be kept dry to reduce the risk of slippage.

Appropriate precautions should be taken to ensure that the cable connecting the Footswitch to the ACM does not create a hazard in the operating room.

## Installing Cylinder Heater Band

### Inspect the Cylinder Heater Band

Inspect the Cylinder Heater Band for any signs of physical damage to the cable and connector.  If physical damage is found or the Cylinder Heater Band does not perform within specification, contact AtriCure for replacement.

### Wrapping the Cylinder Heater Band

Place the Cylinder Heater Band around the bottom of the N2O cylinder, with the cord at the top of the Heater Band, and secure all 6 tensioning spring retainers as shown in Figure 6.  Start with the top and bottom spring retainers.  Be sure heater band is flat against the cylinder (no wrinkles, no puckers).

NOTE:	Cylinder Heater Band should be positioned approximately 2.0 inch (5.0 cm) from the cylinder bottom for optimal performance.

NOTE: If the nitrous oxide cylinder is too large or too small for the supplied ACM Heater Band to fit securely, contact AtriCure Customer Support for replacement.

Figure 6 – Cylinder Heater Band Installation

### Connecting and Disconnecting the Cylinder Heater Band

With the connector alignment arrow in the 12 o’clock position, push the Cylinder Heater Band Connector into the Cylinder Heater Band Receptacle on the rear panel of the ACM as shown in Figure 7.  To disconnect, firmly grasp the connector body, push down on the locking latch, and pull out from the Cylinder Heater Band Receptacle.

NOTE:	Cylinder Heater Band must be connected to the ACM prior to powering up the unit.  However, the ACM unit can operate without the Cylinder Heater Band in place but the system performance will not be optimal.

NOTE:	It is highly recommended that the ACM unit and the cylinder heater band be activated at least 15 minutes prior to first use to allow for the cylinder pressure to reach optimal operating level.

Figure 7 – Connecting the Footswitch, Cylinder Heater Band, and Power Cord to ACM

## Installing N2O Gas Line Hose

### Inspect the N2O Gas Line Hose Module

Inspect the N2O Gas Line Hose for any signs of physical damage to the hose and pneumatic connectors on both ends, including pressure gauge, pressure relief vessel, and the desiccant filter assembly.  If physical damage is found do not use the N2O Gas Line Hose and contact AtriCure for replacement.

NOTE:	Refer to the ACM Tank Hose Subassembly Cart Attachment Instruction (P000862) for details on securing the N2O Gas Line Hose Module to the AtriCure System Cart.

### Connecting and Disconnecting the N2O Gas Line Hose with ACM unit

- Align the quick-connect gas connector from the gas hose line with the ACM N2O inlet port.

- Insert and push in the connector until you hear it “click” in place and the connection is fully seated and secured from unlatching.

- To disconnect the N2O gas line connector from the ACM, slide back on the N2O inlet port connector sleeve and remove.  See Figure 8.

NOTE:	The N2O cylinder valve should be in the CLOSED position during connection and disconnection from the ACM unit.

Figure 8 – Aligning the N2O gas line quick connect to the ACM N2O inlet port

### Connecting and Disconnecting the N2O Gas Line Hose with Gas Cylinder

- Align the CGA326 hand knob (or equivalent cylinder coupler) from the gas hose line with the N2O gas cylinder’s threaded connection port, and screw in place by turning the hand knob in clockwise rotation.

- Hand tighten the connection and verify there is no leak when cylinder valve is opened.

- If a leak is detected, tighten the hand knob connection further.  A wrench may be needed to tighten the connection.

- If the leak cannot be eliminated, either the N2O gas line hose or the N2O gas cylinder may need to be replaced.

- To disconnect from the gas cylinder, turn the hand knob in a counter-clockwise rotation until the connection is disengaged.

NOTE: Keep gas cylinder or gas line canisters standing vertically, in an up-right position

NOTE: Only medical grade N2O gas and an N2O gas cylinder should be used in the gas line connection.

## Installing N2O Exhaust Hose

- Align the 50 ft (15m) exhaust hose with the N2O Exhaust port of the ACM and push in place.

- Run the exhaust hose to an appropriate venting mechanism that meets the hospital and local regulation.

NOTE: The venting mechanism needs to vent at least 60L per minute.

- To remove the exhaust hose from the ACM unit, simply pull back on the tubing until it releases.  See Figure 9.

NOTE:	It is not necessary or required for the N2O Exhaust Hose to be removed from the ACM unit at the end of the procedure.  The hose may be left in place.

Figure 9 – Connecting the N2O gas line hose and N2O exhaust hose to the ACM

## Connecting and Disconnecting the Handpiece

- Connect the Handpiece directly to the ACM.

- Insert the Handpiece pneumatic interface connectors into the pneumatic receptacles on the front panel of the ACM, ensuring that each pneumatic inlet and exhaust connection is fully seated and secured.

- Insert the thermocouple connectors into their respective color coded connection ports on the front panel of the ACM.  See Figure 10.

NOTE:	Typically, you will connect the Handpiece to the ACM when the ACM has been powered up and is in READY operating mode (see Section 4.2 regarding the READY mode).  However, the Handpiece may be connected when powered up, or prior to powering up the ACM.

NOTE:	Once the Handpiece has been connected, it cannot be disconnected from the ACM by pulling on the cable.  To disconnect the Handpiece, slide back on the pneumatic connector sleeve and remove it from the ACM pneumatic receptacle.

NOTE:	Connection and disconnection of the Handpiece with the ACM unit should ONLY take place with the ACM unit in the READY state, N2O cylinder valve in the CLOSED position, and ACM unit fully depressurized by pulling on the System Pressure Relief Knob if the unit has been pressurized prior.

NOTE:	Refer to the Handpiece instruction sheet for more detailed information about connecting the Handpiece to the ACM in a sterile environment.

Figure 10 – Connecting the Handpiece to the ACM

# Instructions For Use

## Powering Up the ACM

- Ensure that the ACM has been plugged into a proper hospital power receptacle.

NOTE:	Do not use extension cords or three-prong to two-prong adapters where applicable.  The power cord assembly should be periodically checked for damaged insulation or connectors.

- Apply power to the ACM unit by flipping the ON/OFF power switch located on the power entry module at the rear panel. The GREEN LED on the power switch should energize if power is being applied to the system.

- The ACM unit will go through a system Power On Self Test (POST).   If POST passes, the system transitions to the READY mode. If any portion of the POST fails, the system transitions to the FAULT mode.

NOTE:	Refer to Section 4.2., below, for a full description of the READY and FAULT modes, as well as all the other operating modes.

- It is highly recommended that the ACM system be prepared and powered up well in advance of start of the procedure to verify its operational readiness as well as preheating of N2O cylinder gas pressure to raise it to its optimal performance level.

NOTE:	N2O cylinder valve must be open for the system to detect its presence and to start the preheating cycle.

NOTE:	Heating of the N2O cylinder is automated and controlled by the ACM unit only as needed and when appropriate.  Operator intervention is not required.

## Operating Modes

The ACM operates in one of five modes:  READY, FREEZE, DEFROST, ERROR and FAULT modes.  These modes are identified by the system status indicator LEDs and the ablation status indicator LEDs located on the front of the ACM unit.  See Figure 11, below.

### READY Mode – This mode is entered automatically upon successful execution of Power On Self Test when the unit is first turned on, or following DEFROST Mode upon venting of the Handpiece by depressing the Ablation Activation Switch or the Footswitch.  This indicates that the system is ready for the next cryo ablation run.

### FREEZE Mode – This Mode is entered from the READY Mode when the user initiates the cryo ablation cycle by pressing and releasing the Ablation Activation Switch or the Footswitch.  In this mode the N2O gas is allowed to cycle through the handpiece causing a temperature drop to take place at the handpiece probe.

### DEFROST Mode – This Mode is entered automatically from FREEZE Mode upon expiration of the ablation timer, or manually by the operator when the Ablation Activation Switch or the Footswitch is actuated while in the FREEZE Mode.  In this mode the handpiece probe temperature is actively forced towards the ambient temperature.

### Once the handpiece temperature has reached 10C or higher, the DEFROST status LED will start to blink to indicate to the user the ACM unit is ready to transition back to the READY Mode or return to the FREEZE Mode to start the next ablation cycle.

### International units:

### The ACM unit returns to the READY Mode by pressing and releasing the Ablation Activation Switch or the Footswitch.  The ACM unit will automatically vent and depressurize the handpiece in preparation for the next freeze cycle.

### US units:

### The ACM unit can be returned to the FREEZE Mode by pressing and releasing the Ablation Activation Switch or depressing and releasing the Footswitch, which causes the handpiece to vent and start the freeze cycle again.

### The ACM unit can be returned to the READY Mode by pressing the Ablation Activation Switch or depressing the Footswitch and holding it in place for 2 seconds or longer.  The ACM unit will automatically vent and depressurize the handpiece.  Once depressurized, the Ablation Activation Switch or the Footswitch can be released.

NOTE:	ACM does allow early transition out from the DEFROST Mode into either the READY Mode or the FREEZE Mode upon user request and upon handpiece temperature reaching above 0C.

NOTE: 	Handpiece probe temperature may drop temporarily during depressurization phase (transition from Defrost to Ready state).

### ERROR Mode – This Mode is entered upon detection of any recoverable error conditions during any Mode excluding the FAULT Mode. The ACM unit will display the corresponding error message when in READY Mode, and may halt operation temporarily upon detection of the error condition.  These are recoverable errors and do not impact continued use of the system.  The System Maintenance LED will be active to alert the user that an error condition has occurred.

### To reset the Maintenance LED, hold down both the Timer Increase and Timer Decrease button and power cycle the unit.  The system will undergo power-on self test for system functionality confirmation and return to the operating state if appropriate.

### FAULT Mode – This Mode is entered upon detection of any unrecoverable error condition during any Mode. The system is inoperable in this Mode until the unit is first power cycled, and only if the fault condition no longer exists or has been remedied.

## Delivering Cryo Energy

### Connect the Handpiece, Pneumatics, Cylinder Heater Band, and Footswitch

Connect the Handpiece, N2O Gas Line Hose, N2O Exhaust Hose, Cylinder Heater Band, and Footswitch as described in Sections 3.4. through 3.8., and note no errors are indicated by the ACM unit.  The ACM should indicate that the system is in the READY mode, and the display showing ablation parameters.  See Figure 11.

If a Handpiece or the Handpiece thermocouple is not connected, the ACM unit will indicate such by displaying an error message in the probe temperature display window with “E-H”.  See Figures 12A and 12B.

Figure 11 – ACM Indicating READY Mode

Figure 12A – ACM Indicating Missing Handpiece Thermocouple

Figure 12B – ACM Indicating Missing Handpiece

### Position the Handpiece

To position the Handpiece, follow the Instructions for Use provided with the Handpiece.

### Delivery of Cryo Energy

- To start the cryo ablation cycle, simply press and release the Ablation Activation Switch or the Footswitch. The ACM unit’s Ablation Status Indicator LED will transition from Ready to Freeze.

- N2O starts to flow through the ACM pneumatics and into the cryo ablation handpiece, and the probe temperature starts to drop.

- Upon probe temperature dropping below the set threshold (value is programmable with default value of -40°C) the ablation timer starts its decrement from a preset value (this value is also programmable with default value of 120 s).

NOTE: 	The Freeze Mode can be manually terminated at any moment by pressing and releasing the Ablation Activation Switch or the Footswitch, causing the ACM unit to transition into Defrost Mode.

NOTE: 	The Ablation Timer value can be manually adjusted from its start value either prior to the start of the ablation cycle or during the freeze cycle using the Timer Increase and Timer Decrease key pad.  Each key press increases or decreases the timer value by 10 seconds.

NOTE: 	The factory default values of 120 seconds for ablation duration and -40°C for the probe temperature threshold value can be altered by the operator, and set as a site default values, following the parameter entry procedure as defined in Section 4.4.

- When the ablation timer reaches zero, the ACM unit automatically transitions into Defrost Mode.  As it transitions into the Defrost Mode, the Ablation Timer display will now show the actual total ablation run time.  The ACM unit’s Ablation Status Indicator LED will transition from Freeze to Defrost.

- At this point, the ACM unit can be made to transition into one of two modes depending on the ACM model:  Freeze Mode or Ready Mode.

International  – ACMs

Defrost to Ready:

Press and release the Ablation Activation Switch or the Footswitch.

The ACM unit vents and depressurizes the handpiece.

The ACM  Ablation Status Indicator LED will transition from Defrost back to Ready.  As it transitions back to the Ready Mode, the Ablation Timer display also resets to the target ablation time value.

Once the Handpiece has been depressurized, it is ready for the next ablation run, which may take place immediately thereafter.

US  – ACMs

Defrost to Freeze:

Press and release the Ablation Activation Switch or the Footswitch.

The ACM  Ablation Status Indicator LED will transition from Defrost back to Freeze and the ablation cycle starts again.  As it transitions back to the Freeze Mode, the Ablation Timer display also resets to the target ablation time value.

Defrost to Ready:

Press and hold the Ablation Activation Switch or the Footswitch for three seconds, until the ACM unit begins to vent and depressurize the handpiece. Total depressurization requires 30 seconds to complete.

The ACM  Ablation Status Indicator LED will transition from Defrost back to Ready.  As it transitions back to the Ready Mode, the Ablation Timer display also resets to the target ablation time value.

Once the Handpiece has been depressurized, it is ready for the next ablation run, which may take place immediately thereafter.

NOTE: 	The transition from Defrost to either Freeze Mode or Ready Mode will not occur until the probe temperature is warmer than 0°C .  Actuation of the Ablation Activation Switch or the Footswitch when the probe temperature is colder than 0°C is simply ignored by the ACM unit.

NOTE: 	The Defrost status LED will start to blink indicating to the operator when the ACM unit is ready to transition back to either the Freeze Mode or the Ready Mode.

NOTE: 	Handpiece probe temperature may drop temporarily during depressurization phase (transition from Defrost to either Freeze Mode or Ready Mode).

NOTE:  Total Handpiece depressurization requires 30 seconds to complete.

### Powering down the ACM

Upon completion of the procedure, the ACM unit can be powered down.

NOTE: 	Verify that the ACM unit is in the Ready Mode prior to powering down the unit.

If the unit is not in the Ready Mode, cycle the ACM by continuously pressing the Ablation Activation Switch or the Footswitch until the unit returns to the Ready Mode.

CAUTION: 	Do not remove cryo handpiece from the  ACM unit if the ACM has been powered down prior to the system being in the Ready Mode as N2O gas may be trapped within the cryo handpiece interface and the handpiece couplers may still be under pressure.

To remove any trapped N2O gas, turn the ACM unit back ON.  Depress the Ablation Activation Switch or the Footswitch to initiate an ablation. After two seconds, stop the ablation by depressing and releasing the Ablation Activation switch or the Footswitch. Depress and hold the Ablation Activation Switch of the Footswitch until depressurization is initiated, then release the Ablation Activation Switch or the Footswitch. Wait 30 seconds for complete depressurization, and then turn OFF the unit utilizing the ON/OFF switch.

When the ACM enters the Ready Mode, the  ACM may now be powered down, and the cryo handpiece can removed from the ACM  safely and disposed of properly.

NOTE: 	When unplugging the handpiece, remove the large diameter tube first.  If a rush of gas is heard or felt, pause before unplugging the small diameter tube until no gas is noticed.

- Close the N2O gas cylinder valve.

- Depressurize the ACM by pulling on the Pressure Relief Valve Knob located in the rear of the ACM. This will remove excess N2O that may be trapped within the ACM pneumatic system.

- Unplug the power cord from the hospital mains outlet, and position on the Cart.

- Disconnect the cryo handpiece from the ACM unit and dispose of properly.

- Position the footswitch from the ACM on the Cart.

- Return ACM Cart back to proper storage.

## Parameter Entry Mode / Adjusting the System Default Parameter Values

- Power up the ACM unit. The system will go through its Power On Self Test (POST) routine.

- Upon completion of POST, press and hold down the Ablation Activation Switch for approximately three seconds and release.

- The User Display will start to blink indicating entry into the parameter entry mode.

- There are two parameters that can be modified by the operator:

- Ablation Timer / Duration: Freeze duration value in seconds

- Probe Temperature Threshold: Temperature value at which the ablation timer starts its countdown sequence

- Ablation Timer / Duration Adjustment

- Using the Timer Increase and the Timer Decrease key pad adjust the timer value until the desired time value is shown.  The displayed value corresponds to the desired length of freeze time in units of seconds. Valid range is from 010 to 300 seconds in 10 second increments.

- Once the desired time value is displayed, press the Ablation Activation Switch to accept the value.

- The system will now start blinking the display of the probe temperature threshold parameter.

- Probe Temperature Threshold

- This value is used by the system to start the ablation timer countdown when the probe temperature has dropped below the set threshold value.

- Using the Timer Increase and the Timer Decrease key pad adjust the probe temperature threshold value until the desired temperature value is shown.  The displayed value corresponds to the desired probe temperature threshold value in units of degrees Celsius. Valid range is from -10°C down to -70°C in 5°C increments.

- Once the desired temperature value is displayed, press the Ablation Activation Switch to accept the value.

- The entry of the parameter value is now complete, and the system exits the parameter entry mode and returns to the operating mode with the new set of parameters.

- Factory default values are:

- Ablation Timer / Duration = 120 seconds

- Probe Temperature Threshold = -40°C

# Troubleshooting

Use the following sections to help troubleshoot possible problems with the ACM.

To clear a fault or error code from the ACM:

- Power OFF the ACM

- Depress and hold both the increase (+) and decrease (-) buttons

- Power ON the ACM. An “E” will be displayed in the N2O gauge display

- Release the increase(+) and decrease (-) buttons simultaneously

- ACM will cycle through its startup test and continue into Ready Mode

- If the Fault/Error Code returns, repeat steps 1-5, ensuring complete depression of the buttons

## No ACM Power / Display Function

If there is no ACM power or display function, attempt to correct this problem using the following checklist.

| Possible Cause | Solution |
| --- | --- |
| ACM not turned on | Turn power on |
| ACM not plugged in | Confirm electrical connections and then turn power on |
| Blown fuse | Replace fuse, confirm electrical connections and turn power on. |
| Internal ACM failure | Contact AtriCure Customer Service |

## ACM Fault Codes

If a fault condition should occur, the N2O Gas Gauge display on the front panel will display the following fault code.  Contact AtriCure Customer Service for support.

NOTE: The ACM will be unusable until the P-Error is resolved.

| DISPLAY MESSAGE | DESCRIPTION | SOLUTION |
| --- | --- | --- |
| P1 | System Controller Board – Fuse F2 blown | Contact AtriCure Customer Service for support |
| P4 | Cold junction temperature out of range | System Controller Board requires service. Contact AtriCure Customer Service for support |

## ACM Error Codes

If an error condition should occur, the N2O Gas Gauge display on the front panel will display the following error code while the ACM unit is in the READY Mode.  The System Maintenance Needed indicator LED will also be lit.  Use the table below to attempt to resolve the following application errors.

NOTE: The ACM will remain usable while displaying an E-Error

| DISPLAY MESSAGE | DESCRIPTION | SOLUTION |
| --- | --- | --- |
| E1 | System Controller Board Error – Fuse blown | Contact AtriCure Customer Service for support |
| E2 | System Controller Board Error – Triac inoperative | Contact AtriCure Customer Service for support |
| E3 | Handpiece pressure regulator out of range | Contact AtriCure Customer Service for support |
| E4 | Internal pressure difference greater than check value - Cylinder over heated (pressure gage over 900psi) - In-line N2O gas filter dirty | Wait for N2O Cylinder to cool down, or Replace cylinder with cool cylinder, or Pull Red Knob to vent pressure, or Visually inspect In-line filter element for debris (see section 9), or Contact AtriCure Customer Service for support |
| E6 | N2O Cylinder Temperature High - N2O Gas Cylinder Empty, or - Heater Band is not attached to tank, or - Heater Band secured improperly - Bottom of Heater Band is not approximately 2 inches from the bottom of the cylinder - Some spring retainers are not secured | Reposition Heater Band, or Secure all 6 tensioning spring retainers, or Wait for N2O Cylinder to cool down, or Replace with new N2O gas cylinder, or Contact AtriCure Customer Service for support |
| E7 | Trapped N2O gas in the Inlet Line - Trying to enter Freeze Mode before connecting the handpiece), or - Powering down ACM before fully depressurizing the handpiece, or - Leaking Inlet Valve Note: User may be unable to connect handpiece to the ACM. | For ACMs WITH an “L” on the top of the unit: Turn off ACM for a minimum of 4 seconds and turn the ACM back on, or If problem persists, contact AtriCure Customer Service for support. |
|  |  | For ACMs WITHOUT an “L” label on the top of the unit: Close the N2O cylinder. Connect the Handpiece Thermocouple Connections. Enter into Freeze Mode. Pull the red System Pressure Relief knob in the back of the unit until all gas has been released. Depress Button to initiate Defrost.Depress Button and hold 3-seconds for Ready state Connect probe and re-open the N2O cylinder, or If problem persists, contact AtriCure Customer Service for support. |
| -System Maintenance Light illuminated (no error code displayed) | No cylinder heater band current Cylinder heater band - Heater Band is not secured to tank (i.e. lying on floor, table, etc.), or - Heating element malfunction, or - Thermal cutoff switch was activated (heater band reached 65C – cylinder will feel very warm to the touch adjacent to the heater band) Note: May notice hot rubber smell Caution: cylinder heater band may be hot to the touch. | Verify Cylinder Heater Band is wrapped around the N2O gas cylinder, and all latches engaged, or Wait for Cylinder to cool down and allow thermal cutoff switch to reset, or Replace Cylinder Heater Band if defective heating element is determined, or Contact AtriCure Customer Service for support |

## Handpiece Error Codes

If Handpiece related error condition should occur, the Probe Temperature display on the front panel will display the following error code.  Use the table below to attempt to resolve the following application errors.

| DISPLAY MESSAGE | DESCRIPTION | SOLUTION |
| --- | --- | --- |
| E-H | READY Mode Only Handpiece Not Connected. | Connect Handpiece Thermocouple Connection |
|  | READY Mode Only Defective Thermocouple. | Replace Handpiece. If Problem Persists, Contact AtriCure Customer Service. |
| E-L | READY Mode Only Defective Thermocouple. | Replace Handpiece. If Problem Repeats, Contact AtriCure Customer Service. |
| HI | FREEZE and DEFROST Mode Only Probe temperature has risen above the operating limit. | Replace Handpiece. If Problem Repeats, Contact AtriCure Customer Service. |
| LO | FREEZE and DEFROST Mode Only Probe temperature has dropped below operating limit. | Replace Handpiece. If Problem Repeats, Contact AtriCure Customer Service. |

# Symbols Used

| Attention: consult accompanying documents |  |
| --- | --- |
| Dangerous Voltage |  |
| Power OFF |  |
| Power ON |  |
| Alternating Current |  |
| Fuses |  |
| Equipotential Terminal |  |
| Type CF Applied Part |  |
| Start |  |
| READY |  |
| FREEZE |  |
| DEFROST |  |
| Gas Gauge (Time Remaining) |  |
| Timer |  |
| Timer Decrease Button |  |
| Timer Increase Button |  |
| Probe Temperature |  |
| Connector |  |
| Gas Inlet |  |
| Gas Outlet |  |
| Temperature Adjust |  |
| Gas Exhaust |  |
| Power |  |
| Maintenance Needed |  |
| Low Gas Pressure |  |
| Cylinder Heater Band |  |
| Footswitch |  |
| Maximum Pressure |  |
| Follow instructions for use |  |

# Technical Specifications

## Mechanical Specifications

Size: 	17.5 in (W)  27.0 in (D)  4.5 in (H) maximum

- 44.5 cm (W)  68.6 cm (D)  11.4 cm (H) maximum

Weight:  	50 lb. (23 kg) absolute maximum.

## Environmental Specifications

Operational temperature:  		+10C to +40C

Storage and Transit temperature:	–35C to +54C

Humidity:		15 to 90% relative humidity

## Electrical Specifications

ACM1:	100-120VAC, (115 VAC nominal), 50/60 Hz

ACM2:	220-240VAC, (230 VAC nominal), 50/60 Hz

## Mains Fuses

100 -120VAC, 50 / 60 Hz,:  Replace fuses as marked: 
4.0A/250V, T-lag, 5  20 mm, UL Recognized, IEC Approved

220-240VAC, 50 / 60 Hz,:  Replace fuses as marked: 
2.0A/250V, T-lag, 5  20 mm, UL Recognized, IEC Approved

## Footswitch Specifications

- Moisture protection rating:	IPX8

## Equipment Type / Classification

- Class 1 Equipment

# Preventive Maintenance and Cleaning of ACM

## Preventive Maintenance

Perform annual preventative maintenance procedures to ensure all ACM components are functioning as defined within this manual.  Pay particular attention to operational and safety features, including but not limited to:

- Electrical power cords for fraying, damage, and proper grounding

- AC power switch

- Indicator damage (Power On, Fault, Maintenance Needed, Low Cylinder Pressure, Cylinder Heater Band Active, Footswitch Active, Ready, Freeze, Defrost)

- N2O Gas Gauge display damage or loss of information

- Ablation Timer display damage or loss of information

- Temperature display damage or loss of information

- Handpiece electronic interface connector damage, cracking or inability to insert and latch Handpiece connector

- Handpiece pneumatic interface connector damage or inability to insert and latch Handpiece pneumatic connector

- Carrying handle damage or inability to fold

- Rubber feet damage, cracking or inability for the ACM to remain stable on a flat surface.

- Rubber alignment cup damage, cracking or inability for the ASB/ASU to remain stable atop ACM and within the alignment cup.

- Footswitch cord fraying or damage

- Footswitch connector damage cracking or inability to insert and latch footswitch plug

- Footswitch pedal damage; check activation by pressing and releasing the pedal

Other medical equipment that may be used simultaneously with the ACM should also be inspected for damage.  Specifically, check for insulation damage of electrical cables and associated connectors.

Visually inspect the footswitch for fluids or other infectious hazards.  Clean as necessary using the instructions in Section 8.2.

The ACM does not have any customer serviceable parts aside from mains fuses, gas line desiccant filter, and Inline gas filter element.  For servicing issues, contact AtriCure, Inc. at:

7555 Innovation Way

Mason, Ohio 45040

USA

Telephone:	513-755-4100

866-349-2342 Toll Free

## Cleaning and Disinfecting

NOTE: Do not spray or pour liquids directly on the unit.

NOTE: The unit and/or accessories cannot be sterilized.

WARNING: Ensure Isopropyl Alcohol (IPA) is completely dry before operating the unit.

CAUTION: Avoid caustic or abrasive cleaners

Guidelines

The following guidelines are recommended for cleaning the unit. It is the user’s responsibility to qualify any deviations from these processing methods.

1. Disconnect the unit or cart from the outlet before cleaning.

2.  If the unit and/or accessories are contaminated with blood or other body fluids, they shall be cleaned before the contamination can dry (within two hours of contamination).

3. The outer surfaces of the unit and/or accessories shall be cleaned with 70% -90% Isopropyl alcohol (IPA) wipes for a minimum of two minutes. Do not allow fluids to enter the chassis.

4. Pay attention to all areas where fluids or soil may gather, such as under/ around the handles or any tight crevices/ grooves.

5. Dry the unit and/or accessories with a dry, white lint-free cloth.

6. Conduct a final confirmation of the cleaning process by visually inspecting the white cloth for remaining soil.

7. If soil remains on the white cloth, repeat steps 3 through 6.

8. After cleaning is complete, turn the unit on to perform Power On Self-Test (POST). If any errors are received, contact AtriCure to begin return process.

# Servicing of ACM Unit

The ACM unit does not have any customer serviceable parts aside from fuses and gas filter.  The manufacturer does not require or recommend routine calibration / output verification.

As stated in Section 4, the ACM performs its own maintenance check every time it is powered ON. The check is called the Power On Self Test (POST).  No calibration is needed for the system and the unit will generate an error if further maintenance is required.

Following identifies servicing that can be performed on the ACM unit based on Section 5, Troubleshooting guide.

## Replacement of AC Mains Fuses

### Tools and Parts

- Needle Nose Pliers

- Fuses:

| ACM Model | Fuse Type | Manufacturer | Part Number | AtriCure P/N |
| --- | --- | --- | --- | --- |
| ACM1 | T 4A L 250V | Schurter | 0034.5049 | C002262 |
| ACM2 | T 2A L 250V | Schurter | 0034.5046 | C002261 |

### Procedure

The ACM unit has been pre-set at the factory to either 110V – 120V (ACM1) or 220V – 240V (ACM2).  The Rating Label below the Power Entry Module on the back panel of the ACM indicates the selected Input Voltage for this unit.  This setting should only be adjusted by the manufacturer or by an authorized AtriCure technical service representative.

NOTE:	ACM unit should be powered off and unplugged before continuing with the fuse replacement procedure.

- Determine the fuse type by looking at the ACM Model Number or the ACM Rating Label.

- Using needle nose pliers, carefully extract the fuse box from the power entry module by squeezing down on the fuse box tabs in the slots as shown.

- Replace the (2) two fuses located in the fuse box. Make sure the fuses are lined up straight.

- Line up the fuse cartridge so the guide tab is towards the power entry side as shown.

1

- Slide the fuse box back into the power entry module and seat it firmly.

- Confirm operational status by plugging in the ACM and powering it up. Check to see that the self-test is completed without errors.

## Inspection and Replacement of In-Line N2O Filter

### Tools and Parts

- 3/32 inch Allen Wrench

- 1.0 inch hex Socket

- Socket Wrench

- Filter:

| In-Line N2O Filter | AtriCure P/N | Type | Manufacturer | Part Number |
| --- | --- | --- | --- | --- |
| Filter | C002458 | TF Series | Swagelok | SS-4F-K4-60 |
| Filter Bonnet Nut | C002500 | TF Series | Swagelok | SS-4TF-K5 |
| Filter Gasket | C002501 | TF Series | Swagelok | SS-4TF-K2 |

### Procedure

NOTE:	ACM unit should be totally depressurized, in the ready state, then powered off and unplugged before continuing with the In-Line N2O Filter Inspection and/ or replacement procedure.

- Turn over the ACM unit so that the bottom side is up.

- Remove the 4 screws holding the filter service cover in its place using the 3/32 inch Allen wrench.

2

3

4

6

- Using the 1 inch (25 mm) hex Socket, loosen the filter bonnet nut (Note: snap ring does not need to be removed).

- Visually inspect the outer surface of the filter element for an accumulation of debris.  If more than half of the outer surface is coated with debris proceed to step 5 to replace the filter element, otherwise insert the Filter Element into the housing, verify the filter gasket is in position and thread the Filter Bonnet Nut tight onto the housing proceed to step 7.

- Remove the old filter element  and filter gasket and replace with new.  Breakdown of each component is shown below for reference.

3

2

1

1. Filter Bonnet Nut

2. Filter Element

3. Filter Gasket

- Insert the filter element centering it in the filter housing.  Insert the filter gasket into the housing.  Install the Filter Bonnet Nut finger tight and tighten to approximately 45.8 ft-lbs (62.2 N-m).

- Replace the filter service cover to the ACM bottom.  Tighten the four cover retaining screws without stripping them.

- Confirm operational status by plugging in the ACM and powering it up. Check to see that the self-test is completed without errors.

## Replacement of Gas Line Desiccant Filter and Tip Washer

### Parts

- Replacement parts:

| Item | Supplied By | Part Number |
| --- | --- | --- |
| Filter Cartridge | AtriCure | F021720 |
| Filter O-ring | AtriCure | F010924 |
| O-ring lubricant | AtriCure | C002502 |
| US Tank Hose Assembly | AtriCure | A000837 |
| International Tank Hose Assembly | AtriCure | A000838 |
| Tip Washer | AtriCure | F021837 |

1

2

3

1. Desiccant Filter Cartridge

2. Filter Housing

3. Tip Washer

### Procedure – Desiccant Filter

- Unscrew the filter cartridge housing by rotating it counter-clockwise.  Refer to the figure below.

Note:  	There are two filter cartridge housing.  Make sure servicing of the filter cartridge is performed only on the unit next to the pressure gauge as shown in the figure below.

6

- Remove the desiccant filter cartridge by rotating it counter-clockwise.  Refer to the figure below.

8

- Replace the desiccant filter cartridge with a new one.

- Replace the filter housing.

### Procedure – Tip Washer (US Model)

- Unscrew the worn out Tip Washer by rotating it counter-clockwise and discard.  Refer to the figure below.

7

- Replace with new tip washer.

## Other Replacement Components

### USA Only

| Item | Supplied By | Part Number |
| --- | --- | --- |
| ACM Footswitch | AtriCure | A000708 |
| N2O Gas Line Hose Assembly (2.5ft / 0.75m) | AtriCure | A000837 |
| N2O Exhaust Hose | AtriCure | C002051 |
| Cylinder Heater Band (CMH15) | AtriCure | A000728 |
| Cylinder Heater Band (CMH22) | AtriCure | A000727 |
| AC Power Cord – 10 ft (USA) | AtriCure | C000262 |

NOTE: Standard cylinder size for the AtriCure Cart is 20lbs with a CGA326 valve.

### International

| Item | Supplied By | Part Number |
| --- | --- | --- |
| ACM Footswitch | AtriCure | A000708 |
| N2O Gas Line Hose Assembly (2.5ft / 0.75m) | AtriCure | A000838 |
| N2O Exhaust Hose | AtriCure | C002051 |
| Cylinder Heater Band (CMH15) | AtriCure | A000728 |
| Cylinder Heater Band (CMH22) | AtriCure | A000727 |
| AC Power Cord – 3.5 m (Europe) | AtriCure | C002090 |

NOTE:  Nitrous Oxide cylinders and valves are country specific.

### Electromagnetic Compatibility

Use of accessories other than identified in Section 9.5.1 and 9.5.2 may result in increased EMISSIONS or decreased IMMUNITY of the AtriCure Cryo Module.

# Accessories

## CMH15, Cryo Module Cylinder Heater Band

Cylinder heater band with dimensions of 12 inch x 15 inch (30.5 cm x 38.1 cm).

## CMH22, Cryo Module Cylinder Heater Band

Cylinder heater band with dimensions of 12 inch x 22 inch (30.5 cm x 55.9 cm).

Spring Extension Kit is available for larger diameter cylinders – contact 
AtriCure customer service

## CMF1, Cryo Module Footswitch

Momentary contact, pedal switch with 15 feet signal cord which can be used in place of the Ablation Activation Switch at the ACM front panel.

# Disposal

Follow local governing ordinances and recycling plans regarding disposal or recycling of device components.

Warranties

Limitation on Liability

This warranty and the rights and obligations hereunder shall be construed under and governed by the laws of the State of Ohio, U.S.A.

AtriCure, Inc. warrants this product to be free from defects in material and workmanship under normal use and preventive maintenance for the respective warranty period shown below. AtriCure’s obligation under this warranty is limited to the repair or replacement, at its option, of any product, or part thereof, which has been returned to AtriCure, Inc. or its Distributor within the applicable time period shown below and which examination disclosed, to AtriCure’s satisfaction, to be defective. This warranty does not apply to any product, or part thereof, that has been: (1) adversely affected due to use with devices manufactured or distributed by parties not authorized by AtriCure, Inc. (2) repaired or altered outside AtriCure’s factory in a way so as to, in AtriCure’s judgment, affect its stability or reliability, (3) subjected to improper use, negligence or accident, or (4) used other than in accordance with the design and use parameters, instructions and guidelines for the product or with functional, operational or environmental standards for similar products generally accepted in the industry.  AtriCure has no control over the operation, inspection, maintenance or use of its products after sale, lease or transfer, and has no control of the selection of Customer’s patients.

AtriCure’s products are warranted for the following periods after shipment to the original purchaser:

AtriCure Cryo Module Unit	One (1) Year

AtriCure Cylinder Heater Band 	One (1) Year

AtriCure Gas Line Hose Assembly………………………………………………	One (1) Year

AtriCure Cryo Footswitch	One (1) Year

Grounded Electrical Cord	One (1) Year

THIS WARRANTY IS IN LIEU OF ALL OTHER WARRANTIES, EXPRESS OR IMPLIED, INCLUDING THE WARRANTIES OR MERCHANTABILITY AND FITNESS FOR A PARTICULAR PURPOSE, AND OF ALL OTHER OBLIGATIONS OR LIABILITIES ON THE PART OF ATRICURE, INC. AND IS A PURCHASER’S EXCLUSIVE REMEDY. IN NO EVENT SHALL ATRICURE, INC. BE LIABLE FOR SPECIAL, INCIDENTAL OR CONSEQUENTIAL DAMAGES INCLUDING, WITHOUT LIMITATION, DAMAGES RESULTING FROM LOSS OF USE, PROFITS, BUSINESS OR GOODWILL.

AtriCure, Inc. neither assumes nor authorizes any other person to assume for it any other liability in connection with the sale or use of any of AtriCure Inc. products. There are no warranties that extend beyond the terms presented unless an extended warranty is purchased before the original warranty expires.  No agent, employee or representative of AtriCure has any authority to change any of the foregoing or assume or bind AtriCure to any additional liability or responsibility.  AtriCure, Inc. reserves the right to make changes to products built and/or sold by them at any time without incurring any obligation to make the same or similar changes on products previously built and/or sold by them.