CRYO1 ACCESSORY KIT INSTALLATION INSTRUCTIONS

Table of Contents

Page

1.	About This Installation	3

2.	Begin Installation Here	4

3.	Pressure Switch Manifold Unit Installation	5

4.	Pressure Regulator Unit Installation	12

5.	Temperature Display Unit Installation	14

6.	Tank Heater Band Installation	16

7.	Verify Accessory Kit Installation	21

8.	Care and Maintenance	21

9.	Return of Product	21

10.	Disclaimer Statements	21

Cryo1 Accessory Kit Installation Instruction

## About This Installation

The Cryo1 Accessory Kit consists of four (4) unit installations at this time. The installation of each unit is required for use of the AtriCure® Cryo1 cryo–ablation probe with the frigitronics® CCS-200 Cardiac Cryosurgical System (CCS–200). This kit must be installed by personnel properly trained on the CCS-200.

The four units should be installed in the following order:

- CPS1 – Cylinder Pressure Switch Manifold

- CPR1 – Cylinder Pressure Regulator

- CPT1 – Cylinder Pressure Temperature Display

- CPB1 – Cylinder Pressure Tank Heater Band

WARNINGS

The Cryo1 accessory kit is for use with the frigitronics CCS–200 Cardiac Cryosurgical System and Cryo1–cryo ablation probe.

The Cryo1 accessory kit is used to maintain appropriate operating pressure levels in the compressed nitrous oxide gas cylinders of the frigitronics CCS–200 Console.

Do not use with gas cylinders other than nitrous oxide.

CAUTIONS

Read all instructions carefully for the Cryo1 accessory kit prior to installing the components. Failure to properly follow instructions may lead to injury and may result in improper functioning of the accessory components.

Use of the Cryo1 accessory kit should be limited to properly trained and qualified medical personnel.

To avoid damage to the device, do not drop or toss any of the Cryo1 accessory components.

Do not use the Cryo1 components if damaged in any way.

Do not remove or install Cryo1 accessory components unless the line and exhaust pressure gauge read "0" psi on the CCS-200 console.

Do not restrict, kink, bend, clamp or otherwise damage nitrous oxide lines.

Getting Help

If you need help with this procedure please contact AtriCure, Inc. Customer Support at

1-877-492-1066.

Amendment Control	Form: ControlDoc1.1

| 1.0 | 1.8.04 | Initial Release | MC |
| --- | --- | --- | --- |

## Begin Installation Here

CCS–200 Shut Down

NOTE:

For additional help on shutting down the system, please refer to the frigitronics CCS-200 Operation and Maintenance Manual.

- Turn the Nitrous Oxide Cylinder Valves clockwise to the “OFF” position.

Turn the system On/Off Valves “ON” and “OFF” repeatedly until all pressure gauges read “0” psi and the Exhaust Pressure Gauge needles touch the silver stop.

- Turn system On/Off Knobs to the “OFF” position.

CAUTION:

Do not remove or install probes unless the line and exhaust pressure gauge read “0” psi.

- If attached, remove probe(s).

- Turn Main Meter switch to the “OFF” position.

- The CCS–200 shut down is complete.

## Cylinder Pressure Switch Manifold Installation

NOTE:

This procedure is applicable only for field upgrade to an existing customer CCS-200 System Console.

New consoles received from Atricure manufacturing will have this module pre-installed.

| Cylinder Pressure Switch Manifold Kit |
| --- |
| Item | Description | Quantity |
| 1. | Pressure Switch Manifold Module | 1 |
| 2. | 1/2” – 3/32” Hex Head Screws | 2 |
| 3. | 9/16” Open-end Wrench | 1 |
| 4. | 11/16” Open-end Wrench | 1 |
| 5. | 20 ft. Exhaust Tube | 1 |

Confirm that all items listed above are in the Cylinder Pressure Switch Manifold Kit.

- If N2O tanks are installed, unstrap the tank(s) and remove from the console.

- Remove the yoke assembly by unscrewing the inlet port.

- Using a set of open end wrenches (9/16” and 11/16”) or with an adjustable open-end wrench and 9/16” open end wrench, hold in place the fixed console nut and loosen the inlet nozzle nut.

- Be certain to keep the console nut in a fixed position as you’re loosening the inlet nozzle nut.

- Remove two (2) left rear console plate screws.

- Using the supplied screw (1/2” – 3/32” Hex Head Screw), thread the rear most mounting screw engaging 3 to 4 full turns

- Slide the mounting bracket of the Cylinder Pressure Switch Manifold module and engage with the screw to temporarily hold the manifold module in place

- Engage the N2O inlet valve and hand tighten

- Insert the second screw to Mounting Point 2 and tighten

- Tighten the rear screw (Mounting Point 1)

- Using a set of open end wrenches (9/16” and 11/16”) or with an adjustable open-end wrench and 9/16” open end wrench, hold in place the fixed console nut and tighten the inlet nozzle nut.  Be certain to keep the console nut in a fixed position as you are tightening the inlet nozzle nut.

- Gas exhaust hose installation.

The National Institute for Occupational Safety and Health’s (NIOSH) current recommended exposure limit for nitrous oxide (N2O) gas is 25 parts per million (ppm) averaged over the period of administration.

- Connect the exhaust tube supplied by AtriCure to the Cylinder Pressure Switch Manifold exhaust nipple located on the bottom of the manifold enclosure as shown.

- Safely vent the other end of the hose in one of the following manners:

- Directly to the outside

- Directly to a Non-Recirculated room exhaust ventilation system.

- Directly to a dedicated system for venting anesthetic gases. The flow capacity of the system shall not be less than 40 liters/min.

CAUTION:

- Only use the exhaust hose provided by AtriCure Inc. No substitutes allowed.

- Do not restrict or kink the exhaust hose in any way.

- The Pressure Switch Manifold Module is now installed.

## Cylinder Pressure Regulator Installation

Amendment Control	Form: ControlDoc1.1

| 1.0 | 1.8.04 | Initial Release | MC |
| --- | --- | --- | --- |

NOTE:

This procedure is applicable only for field upgrade to an existing customer CCS-200 System Console.

New console received from Atricure manufacturing will have this module pre-installed.

| Cylinder Pressure Regulator Kit |
| --- |
| Item | Description | Quantity |
| 1. | Pressure Regulator (with rubber boots) | 1 |
| 2. | Band-Clamp (with screws) | 1 |
| 3. | 5/64” Allen Wrench | 1 |

Confirm that all items listed above are in the Cylinder Pressure Regulator Kit.

- The Cylinder Pressure Regulator will be installed on System II of the CCS–200 console. Refer to figure below to identify System I and System II Probe Socket Supply Connectors.

- Remove the clamp from the Cylinder Pressure Regulator by completely removing the two screws with the provided allen wrench.

- Push the Cylinder Pressure Regulator shaft all the way into the desired Probe Supply Socket Connector. There should be no shaft showing when completely inserted.

- Attach Cylinder Pressure Regulator to Supply Socket Connector with clamp.

  - Align the grooves and capture with clamp

  - Align the two clamp pieces by matching up the larger hole with the smaller hole.

  - Tighten the clamp down with the two screws using the allen wrench.

- The Cylinder Pressure Regulator Unit installation is complete.

NOTE:

The color bands around the System ports are for identification and reference only.  Actual color may vary from Green that is shown in this instruction.

## Cylinder Pressure Temperature Display Installation

| Cylinder Pressure Temperature Display Kit |
| --- |
| Item | Description | Quantity |
| 1. | Temperature Display (with velcro pads) | 1 |
| 2. | Alcohol Wipe Pad | 1 |

NOTE:

The Cylinder Pressure Temperature Display Unit must be installed on the same CCS–200 that the Cylinder Pressure Regulator is on.

- The CCS–200 top console area must be cleaned with the provided isopropyl alcohol wipe before mounting the Temperature Display Unit.

CAUTION:

You only have one opportunity to attach the Cylinder Pressure Temperature Display Unit properly. Confirm the proper placement and alignment before removing the red tape liner from the bottom of the Cylinder Pressure Temperature Display Unit.

- Remove the adhesive backing of the Velcro pad from the Cylinder Pressure Temperature Display Unit.

- Locate the center of the CCS–200 console, approximately in line with the Tissue Temperature input connector locale.

- Align the Cylinder Pressure Temperature Display Unit over the center and in line with the front edge of the CCS–200 console top.

NOTE:

The red and black connectors face the front of the CCS–200 console.

- Once comfortable with the alignment, carefully adhere the unit to the CCS–200 console.

- Apply firm pressure to the Cylinder Pressure Temperature Display Unit to ensure a strong bond.

- The Cylinder Pressure Temperature Display Unit Installation is complete.

## Cylinder Pressure Tank Heater Band Installation

| Cylinder Pressure Tank Heater Band Kit |
| --- |
| Item | Description | Quantity |
| 1. | Heater Band | 1 |
| 2. | Heater Controller Interface Cable | 1 |
| 3. | AC Power Cord | 1 |

Confirm that all items listed above are in the Cylinder Pressure Tank Heater Band Kit.

- Loosen the strap used to secure the N2O cylinder.

- Slide out the N2O cylinder to allow some room behind the cylinder and the back panel of the console.

- Wrap the tank heater band around the cylinder and fasten with affixed spring clips.  Start with the end clips, top and bottom.  Adjust the band position as necessary.

- Connect the remaining spring clips to secure the heater band around the tank cylinder.  Position the nitrous oxide cylinder so that the LED status indicators of the Heater Controller are visible to the CCS–200 Console operator.

- Tighten the tank strap to secure the tank to the console.

NOTE: 	Refer to the CCS–200 Operation and Maintenance Manual for additional information regarding cylinder installation.

- Connect the N2O gas inlet hose and tighten the Hand Knob by turning clockwise.  Ensure the coupler is fully engaged and tightened

- Connect the Heater Controller Interface Cable between the pressure switch manifold and the heater controller box as shown below.

- Insert the heater power cord into the Heater Controller power input receptacle and connect to main power line.  Ensure that the Heater Controller power switch is in the OFF position prior to interfacing with the main power line.

- Open the tank cylinder valve by turning the Hand Knob counter-clockwise fully, and verify that the pressure reading is registered on the pressure gauge.

- Apply power to the Cylinder Pressure Tank Heater Band controller.  The Green LED embedded in the power switch should turn on if the AC line power is active, as shown below.  The White LED should illuminate indicating that the heating element of the Cylinder Pressure Tank Heater Band is now active.  The Green LED on the controller box will start to blink with the applied power if there is a communication error between the Cylinder Pressure Heater Controller and the Cylinder Pressure Manifold Module.  Check the Heater Controller Interface Cable connections.

- Controller Box status LED and definition.

| LED Color | LED Condition | Definition |
| --- | --- | --- |
| White | Continuous ON state | Heater Band’s heating element is active. Tank is being heated. |
| Green | Continuous ON state | Tank pressure is within the nominal operating range. |
| Green | Blinking state | Communication error between the Cylinder Pressure Heater Controller and the Cylinder Pressure Manifold Module. Heater Controller Interface Cable may not be connected, fully seated, or defective. |
| Amber | Continuous ON state | Tank is low on N2O; approaching empty state. Replace with new tank. |

- System is now ready for operation

## Verify Accessory Kit Installation

Amendment Control	Form: ControlDoc1.1

| 1.0 | 1.8.04 | Initial Release | MC |
| --- | --- | --- | --- |

WARNING:

Installation of all four (4) of the Cryo1 Accessory Components listed in this document is required to the CCS–200 prior to using the AtriCure Cryo1 cryo–ablation probe.

## Care and Maintenance

  - The CCS-200, including Cryo1 accessory components, should be handled with the care appropriate to any precision medical instrument. Avoid banging, dropping or mishandling the system. Periodically inspect the hoses and fittings for any damage or wear. Other than maintaining an adequate supply of medical grade nitrous oxide, the system is essentially maintenance free.

  - Cleaning

The control console may be wiped down with any mild soap solution that is commonly used to clean equipment in the operating room. Care should be exercised to use only a lightly dampened sponge when wiping down any electronic modules such as the timers and temperature meters. They should be wiped dry after cleaning. Under no circumstances should solution be allowed to enter internally.

## Return of Product

Amendment Control	Form: ControlDoc1.1

| 1.0 | 1.8.04 | Initial Release | MC |
| --- | --- | --- | --- |

If for any reason this product must be returned to AtriCure, Inc., a return goods authorization (RGA) number is required from AtriCure, Inc., prior to shipping.

CAUTION:

It is the responsibility of the health care institution to adequately prepare and identify the product for shipment.

If the product has been in contact with blood or body fluids, it must be thoroughly cleaned and disinfected before packing. It should be shipped in either the original carton or an equivalent carton, to prevent damage during shipment; and it should be properly labeled with an RGA number and an indication of the bio-hazardous nature of the contents of shipment.

Instructions for cleaning and materials, including appropriate shipping containers, proper labeling, and an RGA number may be obtained from AtriCure, Inc.

## Disclaimer Statements

Amendment Control	Form: ControlDoc1.1

| 1.0 | 1.8.04 | Initial Release | MC |
| --- | --- | --- | --- |

Users must assume responsibility for approving the condition of the installation before it is used. AtriCure, Inc., cannot be held liable for any consequential damage, personal injury or damage to property caused by improperly installed products.

This document describes the proper procedures for the product installation. Any deviation from these procedures, which may compromise the function of the product, is the responsibility of the user.

# Text boxes

Mounting

Point 1

Mounting Point 2

Hand Tighten

Inlet Nut

MP #1

MP #2

Engage this end

From Console

Supplied

Secure Clamp with Screws

AtriCure, Inc.

6033 Schumacher Park Drive

West Chester, Ohio 45069     USA

Telephone: 513-755-4100

Technical Support: 877-492-1066

P000561  Rev B

Mounting

Point 2

Inlet

Nozzle Nut

Fixed

Console Nut

Tank Valve Open

Tank Valve Closed

Tighten nut

Hold in place

Inlet Screw

Hold Fixed

Inlet Nut

Console Nut

Loosen

Hold Fixed

Hand Knob

Heater Controller Interface cable

Bottom view of the Heater Controller

Exhaust Nipple

AC Power

Active

Blinking Green LED

Communication

Error

White LED

Heating Element

Active